# Age-Based Individualization of VT Detection Thresholds in Wearable Cardioverter-Defibrillators: A Case Against the Universal 150 bpm Cutoff

Running Title:

DOI:

Jakub K. Rokicki1,2, Andrzej Cacko1 and Marcin Grabowski2

1Department of Medical Informatics and Telemedicine, Medical University of Warsaw, Warsaw, Poland

2First Chair and Department of Cardiology, Medical University of Warsaw, Warsaw, Poland

Manuscript received XXX. Final version accepted XXX.

Address correspondence to: Jakub K. Rokicki, Litewska 14/16, Warszawa, Mazowieckie, 00-581, Poland. Tel.: +48-221169244. Email: jakub.rokicki@wum.edu.pl.

## KEYWORDS

Cardiac defibrillator, sudden cardiac death prevention, ventricular tachycardia, wearable cardioverter-defibrillator.

## Authorship statement

J.K. Rokicki conceived the idea for this letter, conducted the literature review, drafted the manuscript, and is the corresponding author. A. Cacko and M. Grabowski critically reviewed the manuscript for intellectual content and provided substantive revisions. All authors approved the final version of the manuscript and agree to be accountable for all aspects of the work.

## Original submission declaration

The authors confirm that this manuscript has not been previously published, is not currently under consideration for publication elsewhere, and will not be submitted to another journal while under review at the Journal of Innovations in Cardiac Rhythm Management.

## Conflict of interest

Jakub Rokicki and Marcin Grabowski declare speaker’s honoraria from ZOLL. Andrzej Cacko reports no conflicts of interest relevant to this correspondence. All authors confirm no conflict of interest involving Kestra Medical Technologies.

## Funding

No funding was received for the preparation of this manuscript.

Dear Editor,

We read with great interest the article by Ringquist et al.1 examining heart rate threshold settings for ventricular tachycardia (VT) detection in patients with wearable cardioverter-defibrillators (WCDs). Based on a retrospective, cases-only analysis of 676 appropriate shocks from the ZOLL Medical LifeVest™ registry, the authors suggest that a programmed VT detection threshold of 150 beats per minute (bpm) “should be considered” for WCD recipients. While we appreciate the effort to characterize slow VT in the WCD population and acknowledge that the authors’ language is permissive rather than prescriptive, we are concerned that the conclusion may nonetheless be interpreted as endorsing a fixed 150 bpm threshold in routine practice. We believe the study design does not adequately support this suggestion and that broad adoption risks increasing inappropriate therapy delivery - particularly in younger and more physically active patients.

Our central concern is physiological. The age-predicted maximum heart rate, commonly estimated by the formula 220 minus age (with the regression-derived formula 208 − 0.7 × age, as proposed by Tanaka et al.2), defines the ceiling of achievable sinus rates during exertion. While both formulas carry a standard deviation of approximately ±10 bpm and neither accounts for chronotropic incompetence, they remain useful population-level anchors for programming decisions. The authors’ own data reveal that patients with slow VT (VT150–170) were significantly older than those with fast VT (71 ± 12 vs. 67 ± 11 years; P ≤ 0.01). Critically, for a 71-year-old patient, the formula 220 − age yields an age-predicted maximum heart rate of 149 bpm - virtually identical to the proposed 150 bpm threshold. This means the 150 bpm cutoff “worked” in this older cohort precisely because these patients were physiologically unlikely to reach that rate during sinus tachycardia. By contrast, a 50-year-old patient has an expected maximum heart rate of 170 bpm and a 40-year-old of 180 bpm. For these patients, a 150 bpm threshold would capture routine exercise-induced sinus tachycardia within the VT detection zone, inevitably triggering false alarms or inappropriate shock sequences.

The study design has several limitations that preclude the broad recommendation offered in the conclusion. First, this is a cases-only analysis of appropriate shocks - there is no control group, no data on the total number of WCD recipients during the study period, and, most critically, no data on inappropriate shocks or false alarms. Without the denominator, the finding that 15% of appropriate shocks occurred for VT150–170 is uninterpretable from a programming-optimization standpoint: one cannot recommend a lower detection threshold without quantifying the false-positive cost that threshold incurs. The authors acknowledge this limitation but nonetheless proceed to a broad suggestion for clinical consideration. Second, the absence of inappropriate shock data is particularly problematic given that the very reference the authors cite to support the 170 bpm ICD threshold (MADIT-RIT3) demonstrated its benefit precisely by reducing inappropriate therapy - the exact outcome this study does not measure.

Moreover, the clinical outcome data presented in the article may paradoxically argue against aggressive intervention for slow VT. Among the VT150–170 group, 70% of patients did not experience loss of consciousness (LOC), compared to 50% in the VT>170 group (P ≤ 0.01). Response button use was significantly higher in the slow-VT cohort (36% vs. 24%; P = 0.01), indicating that these patients were conscious, alert, and hemodynamically stable enough to interact with the device. These findings suggest that slow VT in WCD patients is generally better tolerated, which argues for longer observation windows or conditional intervention rather than reflexive shock delivery at 150 bpm. The 60-second VT detection delay already provides a safety margin, but this could be further leveraged with individualized rate thresholds.

The broader evidence from ICD programming trials unequivocally favors higher, not lower, rate cutoffs. The MADIT-RIT trial3 demonstrated that programming ICD therapies for tachyarrhythmias of ≥200 bpm resulted in a 79% reduction in inappropriate therapy and a 55% reduction in all-cause mortality compared with conventional programming (VT zone ≥170 bpm). Consequently, in the same trial delayed intervention strategy (60-second delay at 170-199 bpm, a 12-second delay at 200-249 bpm, and a 2.5-second delay at 250 bpm or higher) showed superiority in terms of inappropriate shock delivery (risk reduction of 76%) with reduction in all-cause mortality of 44%, which missed the statistical significance, compared to conventional programming involving delivering therapies after 2.5 seconds starting from detection of VT at rate above 170 bpm. Detection cutoff described by Ringquist et al.1 was not considered in MADIT-RIT trial on programming ICD therapies in primary prevention patients. The ADVANCE III trial4 showed a 48% reduction in overall therapies with prolonged detection intervals, which was also confirmed in reduction of appropriate shock therapy (51% reduction) and inappropriate shock delivery (22% missing statistical significance). These findings underpin the 2015 HRS/EHRA/APHRS/SOLAECE expert consensus statement (updated in 2019),5 which discusses evidence supporting rate cutoffs for primary prevention patients (“with no VT history”) from 185 up to 200 bpm and recommends a minimum detection duration of 6–12 seconds or 30 intervals, with manufacturer-specific programming recommendations detailed in its Appendix B. While WCDs differ from ICDs - most notably in lacking antitachycardia pacing and offering patient-override capability - the principle that higher detection thresholds reduce inappropriate therapy without compromising safety is directly transferable. Notably, the ASSURE® WCD, the other commercially available device in the United States, already uses a default VT detection rate of 170 bpm,6 suggesting that engineering, regulatory, and clinical considerations led an independent manufacturer to a higher default - though we acknowledge that device defaults reflect multiple factors beyond purely scientific assessment.

We propose an individualized, physiology-guided approach to WCD programming. The conditional VT intervention zone should be set at or near the patient’s age-predicted maximum heart rate (using either 220 − age or 208 − 0.7 × age as a starting estimate, adjusted for individual patient factors), or, in patients with documented VT, below the known VT rate (typically 10–20 bpm below the documented VT rate, consistent with established ICD programming practices5). We acknowledge that the 220 − age formula is an approximation and that individual patients - particularly those with heart failure, a group overrepresented in the VT150–170 cohort (10% vs. 2%; P ≤ 0.001) - may have significantly reduced exercise capacity and chronotropic incompetence, rendering their true achievable sinus rates lower than the formula predicts. However, this very heterogeneity reinforces rather than undermines the case for individualization. A fixed 150 bpm threshold treats all patients identically despite widely divergent physiological profiles, whereas an age-guided starting point, adjusted downward when clinical context warrants (e.g., known chronotropic incompetence, severe heart failure, or beta-blocker therapy), offers a rational framework that can accommodate the full spectrum of patients. Critically, even accepting the limitations of the formula, the risk-benefit balance clearly disfavors a universal 150 bpm cutoff: the marginal benefit of detecting the small proportion of slow VT events must be weighed against the substantially greater risk of inappropriate detections, alarm fatigue, and consequent device non-compliance across the much larger population of WCD recipients who never experience VT.

Compliance is the Achilles’ heel of WCD therapy, as the VEST trial demonstrated.7 Unlike ICDs, WCDs depend entirely on the patient’s willingness to wear the device continuously. Every false alarm degrades trust. The Ringquist et al.1 data show that more than one-third of slow-VT patients pressed the response button, meaning they were conscious and actively engaging with the device alarm. In a population already prone to compliance challenges, programming a VT threshold that captures physiological heart rates will amplify alarm fatigue and may paradoxically increase the risk of sudden cardiac death by driving patients to remove the device altogether.

Finally, we note that two of the three authors are employees of ZOLL Medical, the manufacturer of the LifeVest™ WCD, which uses 150 bpm as its default VT detection rate. While we do not question the authors’ scientific integrity, the reader should consider that the study’s conclusion effectively endorses the manufacturer’s own default programming. Independent, prospective studies comparing age-adjusted thresholds with fixed cutoffs - ideally with inappropriate therapy as a primary endpoint - are needed to definitively inform this programming decision.

In conclusion, while characterizing slow VT in WCD recipients is a valuable contribution, the leap from a retrospective cases-only analysis to suggesting a 150 bpm threshold for broad clinical consideration is not supported by the data presented. The age-dependent overlap between sinus tachycardia and the proposed detection zone, the absence of inappropriate therapy data, and the favorable hemodynamic profile of slow VT all argue against this approach. More broadly, the prevention of sudden cardiac death - whether by ICD, WCD, or any other modality - demands individualization of therapy. A one-size-fits-all detection threshold contradicts the foundational principle that device programming should be tailored to the patient’s arrhythmia substrate, hemodynamic status, age, and functional capacity. The history of ICD programming teaches us that lowering detection thresholds in pursuit of sensitivity comes at an unacceptable cost in specificity, patient harm, and ultimately mortality.3 We urge the electrophysiology community to resist the appeal of a simple universal number and instead adopt age-predicted maximum heart rate as a rational physiological anchor for individualized WCD programming. Prospective, manufacturer-independent studies comparing age-adjusted thresholds with fixed cutoffs - with inappropriate therapy and device compliance as co-primary endpoints - are urgently needed to resolve this question.

References

- Ringquist S, Bianco NR, Tung P. Heart rate threshold settings for ventricular tachycardia in patients with wearable cardioverter-defibrillators. J Innov Card Rhythm Manag. 2026;17(2):6627–6629. [CrossRef] [PubMed]

- Tanaka H, Monahan KD, Seals DR. Age-predicted maximal heart rate revisited. J Am Coll Cardiol. 2001;37(1):153–156. [CrossRef] [PubMed]

- Moss AJ, Schuger C, Beck CA, et al. Reduction in inappropriate therapy and mortality through ICD programming. N Engl J Med. 2012;367(24):2275–2283. [CrossRef] [PubMed]

- Gasparini M, Lunati MG, Proclemer A, et al. Long detection programming in single-chamber defibrillators reduces unnecessary therapies and mortality: the ADVANCE III trial. JACC Clin Electrophysiol. 2017;3(11):1275–1282. [CrossRef] [PubMed]

- Stiles MK, Fauchier L, Morillo CA, Wilkoff BL; ESC Scientific Document Group. 2019 HRS/EHRA/APHRS/LAHRS focused update to 2015 expert consensus statement on optimal implantable cardioverter-defibrillator programming and testing. Europace. 2019;21(9):1442–1443. [CrossRef] [PubMed]

- Aldaas OM, Birgersdotter-Green U. Advancements in automated external and wearable cardiac defibrillators. Curr Opin Cardiol. 2025;40(1):15–21. [CrossRef] [PubMed]

- Olgin JE, Pletcher MJ, Vittinghoff E, et al. Wearable cardioverter-defibrillator after myocardial infarction. N Engl J Med. 2018;379(13):1205–1215. [CrossRef] [PubMed]